{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["Sheet1","Sheet2","Sheet3"],"not_rendered":{"charts":0,"images":1,"macros":1,"external_links":0,"defined_names":5,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Sheet1","index":0,"n_rows":46,"n_cols":9,"rows_kept":46,"cut":false,"hidden":false,"merged":[[18,1,18,4],[9,1,9,8],[10,0,10,9],[19,7,20,7],[19,8,20,8],[14,2,14,4],[19,4,20,5],[19,6,20,6],[45,2,45,5],[42,2,42,5],[0,0,8,9],[45,0,45,1],[39,0,39,9],[13,0,13,1],[12,0,12,1],[42,0,42,1],[13,2,13,4],[12,2,12,4],[19,1,20,1],[19,2,20,3],[30,2,30,4],[40,0,40,9],[31,2,31,4],[36,1,38,8],[26,2,26,3],[26,4,26,5],[22,4,22,5],[23,4,23,5],[24,4,24,5],[25,2,25,3],[21,2,21,3],[22,2,22,3],[23,2,23,3],[24,2,24,3],[25,4,25,5],[21,4,21,5]]}
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null,"120-64 Constant Speed Blender 115v"]
["Calibration Report"]
[]
["Sales/ Work Order:",null,"535022",null,null,null,null,"Calibration Date:","2024-07-01"]
["Serial Number:",null,"5350220000002",null,null,null,null,"Recommended Calibration Date:",{"$f":"IF(I13=\"\",\"\",I13+365)","v":"2025-07-01"}]
[null,"Blender SN:","na"]
[]
[]
[]
[null,"Rotational Speed Test:"]
[null,"Target    Speed","Low Point                    Actual Speed",null,"High Point                    Actual Speed",null,"Accepted Deviation","Actual Deviation","Pass / Fail"]
[]
[null,500,null,null,null,null,10,{"$f":"MAX(ABS(B22-C22),ABS(B22-E22),ABS(C22-E22))","v":500},{"$f":"IF(OR(C22=\"\",E22=\"\"),\"Fail\",IF(H22>G22,\"Fail\",\"Pass\"))","v":"Fail"}]
[null,1500,null,null,null,null,30,{"$f":"MAX(ABS(B23-C23),ABS(B23-E23),ABS(C23-E23))","v":1500},{"$f":"IF(OR(C23=\"\",E23=\"\"),\"Fail\",IF(H23>G23,\"Fail\",\"Pass\"))","v":"Fail"}]
[null,2000,null,null,null,null,40,2000,{"$f":"IF(OR(C24=\"\",E24=\"\"),\"Fail\",IF(H24>G24,\"Fail\",\"Pass\"))","v":"Fail"}]
[null,4000,3994,null,4012,null,100,{"$f":"MAX(ABS(B25-C25),ABS(B25-E25),ABS(C25-E25))","v":18},"Pass"]
[null,10000,null,null,null,null,250,10000,"Fail"]
[null,12000,11992,null,12013,null,250,21,"Pass"]
[]
[null,"Tachometer Calibration Information:",null,null,null,null,"117"]
[]
[null,"Model Number:",{"$f":"Sheet2!B2","v":"461920"},null,null,null,null,"Calibration Due Date:",{"$f":"DATEVALUE(Table_ExternalData_1[DueDate])","v":"2024-09-13"}]
[null,"Serial Number:",{"$f":"Sheet2!C2","v":"10055320"}]
[]
[]
[]
[]
[null,"This report meets or exceeds the requirements outlined in API RP 10B-2 Recommended Practice for Testing Well Cements for calibration of laboratory  devices, using certified master test equipment that is traceable to NIST."]
[]
[]
[]
[]
[]
["Technician Tested:",null,"Elvin Le",null,null,null,null,"Date:","2024-07-01"]
[]
[]
["Approved By:",null,null,null,null,null,null,"Date:"]

{"sheet":"Sheet2","index":1,"n_rows":2,"n_cols":4,"rows_kept":2,"cut":false,"hidden":true,"merged":[]}
["AliasID","Model","SerialNumber","DueDate"]
[117,"461920","10055320","2024-09-13"]

{"sheet":"Sheet3","index":2,"n_rows":14,"n_cols":1,"rows_kept":14,"cut":false,"hidden":true,"merged":[]}
[]
[]
["120-60 Constant Speed Blender 115v"]
["120-60-1 Constant Speed Blender 230v"]
["120-60-F Constant Speed Blender 115v"]
["120-60-1-F Constant Speed Blender 230v"]
["120-64 Constant Speed Blender 115v"]
["120-64-1 Constant Speed Blender 230v"]
["120-64-F Constant Speed Blender 115v"]
["120-64-1-F Constant Speed Blender 230v"]
["120-65 Constant Speed Blender 115v"]
["120-65-1 Constant Speed Blender 230v"]
["120-65-F Constant Speed Blender 115v"]
["120-65-1-F Constant Speed Blender 230v"]